[image: image1.jpg]



T.E.N.S., Analog
	Dual channel, modulation/burst/standard modes, complete with accessories and carrying case.
* Federal Law (USA) restricts this device to sale by or on the order of a physician.

Designed to meet the requirement of HCPCS code E0720
TECHNICAL SPECIFICATIONS

 

Channels                              : Dual

Pulse Amplitude                  : Adjustable, 80 mA peak maximum for positive pulse or -20 mA peak maximum for negative pulse into 500 ohm load each channel. 

Pulse Rate                            : (Adjustable) 2 Hz – 120 Hz

Pulse Width                          : (Adjustable) 40 µ S – 260 µ S

Modulation Mode                 : Pulse width is automatically varied in a cyclic pattern over an interval of nominally 5.0 seconds maximum. Pulse width decreases linearly over a period of 1.5 seconds from the control setting value to a value which is decreased 45% maximum. The narrower pulse width will continue for 2 seconds maximum, and then increase linearly over a 1.5 second period to its original value. The cycle is then repeated. 

Burst Mode                           : Bursts consist of 9 ±2 pulses at maximum pulse rate. 

                                                 Bursts occur twice every second. 

Wave Form                           : Asymmetrical Bi-Phasic square pulse. 

All of the adjustable parameters are setting at the midpoint of the specified range. 

Max. Charge per pulse       : 21 micro coulombs maximum. Positive pulse and negative pulse separate. 

Battery                                    : 9 volt battery (alkaline, or nickel-cadmium rechargeable). 

Size                                        : 23x 64x 91mm

Weight                                   : 130 grams (battery included)

There may be up to a ±20% tolerance of all specifications. 


[image: image2.png]



